## Slide 1: Quick and Reliable Method for Cleanup of All 209 PCBs in One Fraction in Environmental Samples
Fluid Management Systems
Watertown MA
Philip Bassignani, Ruud Addink, Philip Germansderfer
Fluid Management Systems
580 Pleasant Street, Watertown, MA 02472 , USA
www.fms-inc.com

## Slide 2: Introduction (1)
POPs (PCDD/Fs, PCBs) continue to attract interest around the world due to strict regulations enforced in many countries
Rapid and quality sample clean up and analysis is needed for many laboratories processing samples
Processing times and cost are important considerations
In the US, EPA methods SW-846, 1668C and 8082A are used for PCBs

## Slide 3: Introduction (2)
Sample extracts in DCM, hexane or toluene
Cleanup for analysis of all 209 PCBs in common in North America - extracts are often in toluene after Soxhlet Extraction
Dioxins and furans can also be run

## Slide 4: Challenges of POPs Sample Prep
Labor intensive, prone to error
Compliance with regulatory procedures and accreditation (lengthy method validation)
Strict QA/QC requirements
Sample matrix complexity
Native background and interferences (can be orders of magnitude higher than analytes)
Pico/femto-gram analyses require ultra pure extract and excellent instrument sensitivity

## Slide 5: Automated Sample Prep
Advantages of Automated Sample Prep
  Rapid Turn Around Time: 30 to 45 Minutes for 6 Samples
  Cleaner Background Interferences: Closed Loop System
  Quality Results: Certified Pre-packaged Columns
  Green Technology: Lower solvent and power use
  QA/QC & Accreditation Requirements: Easier to Manage
  Computerized Method: Instrumentation based prep

## Slide 6: Manual Sample Prep
Advantages of Manual Sample Prep
  Most labs use a Manual Methods for the following reasons:
    No electronics or mechanical components to fail
    No down time due to the system failure
    No service contract
    No capital equipment cost

## Slide 7: Semi-Automated System
Simple to run, no computerized instrumentation
Fast: 30 to 45 min
Closed loop system to give a clean background, low level detection
Use certified pre-packaged columns
Green technology, only vacuum pump uses power
Low solvents
Economical column kits, choice of low fat and high fat column kits
No capital equipment cost
No electronics or mechanical equipment to fail
No downtime
Specification:

## Slide 8: Characteristics of Semi-Automated System (EZPrep)
Small Solvent Reservoirs
Large Solvent Reservoirs
Turn Table
Stage 1 Manifold (waste)
3 Way Valve
Stage 2 Manifold (collection)
Large Solvent
Reservoir 100/180 mL
Small Solvent Reservoirs (40 mL)

## Slide 9
Carbon
Alumina 4g
Classical Plus Acid
High Capacity Acid
Classical Acid
Alumina 6g
XL Capacity
Acid
Columns (1)

## Slide 10
Classical Kit
Classical Plus Kit
Mini Kit
High Capacity Kit
XL Capacity Kit
Columns (2)

## Slide 11
Stage 1 Manifold
Vacuum Gauge
Vacuum Relief Valve
Stopcock
Fitting to 3 Way Valve

## Slide 12
Stage 2 Manifold

## Slide 13: Stage 1: to waste
Vacuum pump
Waste
Rotatory work station
Solvents
Column
assembly
Vacuum gauge

## Slide 14: EZPrep Stage 1

## Slide 15: Stage 2: collect
Manifold with collection tubes
Columns ready for elution

## Slide 16: EZPrep Stage 2 Fractions

## Slide 17: Attributes
  Closed loop system:
    Eliminates background contaminants
    No washing needed.
    Capped solvent reservoirs
  Optimized for solvent reduction while obtaining highest possible recoveries
  Easy sample loading on top of silica column via syringe vial
  Columns connect easy with SNAP connections

## Slide 18: Extracts in hexane - PCBs
  Stage 1: Connect High Capacity Acid Silica and Alumina (no Carbon) and condition with 60 mL of hexane (vacuum, waste)
  Stage 2: Load sample (in hexane, collect Fraction # 1), rinse loading vials with hexane, elute with 160 mL hexane (collect Fraction # 1), remove acid silica, elute alumina with 50 mL dichloromethane (collect Fraction # 1)
  All 209 PCBs are now in Fraction # 1

## Slide 19: SuperVap 6 Concentrator 250 mLs

## Slide 20: SuperVap Concentration/Evaporation
System pre-heated to 55-60 ºC.
Samples evaporated at stable T under 5-6 psi nitrogen.
1 mL extract vial transferred to GC vial (can have direct-to-vial feature).
Recovery standards added (nonane/dodecane).
Extract taken to 10 uL volume with a gentle stream of nitrogen at ambient temperature.

## Slide 21: SuperVap 24 position GC vial Concentrator

## Slide 22: Direct-to-Vial
GC vial

## Slide 23: DFS HRGC/HRMS

## Slide 24: Native PCBs in Serum extract

## Slide 25: Native PCBs in Soil extract

## Slide 26: Extracts in toluene -PCBs
  Stage 1: Connect High Capacity Acid Silica and Alumina (no Carbon) and condition with 60 mL of hexane (vacuum, waste)
  Stage 2: Load sample (in 2-10 mL toluene, collect Fraction # 1), rinse loading vials with hexane, elute with 60 mL hexane (collect Fraction # 1), remove acid silica, elute alumina with 50 mL dichloromethane (collect Fraction # 1)
  All 209 PCBs are now in Fraction # 1

## Slide 27: 13C recoveries PCBs soil
Chart: Recoveries (%)
| PCB_28 | PCB_52 | PCB_77 | PCB_81 | PCB_101 | PCB_105 | PCB_114 | PCB_118 | PCB_123 | PCB_126 | PCB_138 | PCB_153 | PCB_156 | PCB_157 | PCB_167 | PCB_169 | PCB_170 | PCB_180 | PCB_189 | PCB_209
 | 82.5 | 81.1667 | 82.1667 | 80.1667 | 79.3333 | 84 | 79.1667 | 82.3333 | 84 | 82.8333 | 86.8333 | 87.8333 | 86.1667 | 84 | 82.8333 | 78.8333 | 75.5 | 79.3333 | 83.5 | 86
10 g soil in toluene, n=6

## Slide 28: 13C recoveries PCBs salmon
Chart: Recoveries (%)
| PCB_28 | PCB_52 | PCB_77 | PCB_81 | PCB_101 | PCB_105 | PCB_118 | PCB_123 | PCB_126 | PCB_138 | PCB_153 | PCB_156 | PCB_157 | PCB_167 | PCB_169 | PCB_170 | PCB_180 | PCB_209
 | 84.6667 | 85.3333 | 86.8 | 87 | 88.3333 | 89.6 | 83.2 | 73.1667 | 75.6 | 86.6667 | 88.3333 | 80.6667 | 80 | 88.1667 | 81.3333 | 77.5 | 94.4 | 88
2 g salmon in toluene, n=6

## Slide 29: Column Kits with various fat removal capacities for samples in hexane
 |  | STAGE 1 |  |  | STAGE 2
 |  |  |  |  | PCBs
 | Fat Removal | Hexane | Hexane | Hexane | DCM
Column kits | Capacity | conditioning (mL) | sample volume (mL) | Elute Silica (mL) | Alumina (mL)
Classical Plus | 1.0 g | 20 | 30 | 100 | 50
High Capacity | 2.5 g | 40 | 30 | 160 | 50
Extra high Capacity | 5.0 g | 60 | 30 | 180 | 50

## Slide 30: Conclusions (1)
  Samples in toluene (environmental, food): 2-10 mL toluene, collect all 209 PCBs in one fraction using hexane followed by DCM
  Reduced hexane volume needed for silica column because of presence toluene
  Works also for samples in hexane but more hexane needed in that case for silica elution (“toluene effect” not present)
  EZPrep suitable for environmental and food analyses in toluene as solvent. Also suitable for samples in hexane
  High sample throughput → 18 samples/hour
  6 samples in parallel per station
  3 stations fit in one hood

## Slide 31: Conclusions (2)
  System gives excellent recoveries for PCBs comparable to automated systems
  Use of certified pre-packaged columns guarantees low native background
  No worries about breakdown or downtime
  No washing needed
  No cross-contamination
  Low cost

## Slide 32: Questions
Contact us at:
raddink@fms-inc.com (Ruud Addink)
thall@fms-inc.com (Tom Hall)