## Slide 1: Optimized Workflow of POPs Analysis in Environmental and Food Matrices Using Semi-Automated Cleanup and GC-MS/MS
Philip Bassignani, Ruud Addink, Philip Germansderfer
Fluid Management Systems
580 Pleasant Street, Watertown, MA 02472 , USA
www.fms-inc.com
Ruud Addink (FMS) and Dale Walker (Agilent Technologies)

## Slide 2: Overview
Introduction
Automated Hyphenated System Configurations for different Matrices and Extractions
Semi Automated Systems
GC/MS/MS Data

## Slide 3: Challenges of POPs Sample Prep
Labor intensive, prone to error
Compliance with regulatory procedures and accreditation (lengthy method validation)
Strict QA/QC requirements
Sample matrix complexity
Native background and interferences (can be orders of magnitude higher than analytes)
Pico/femto-gram analyses require ultra pure extract and excellent instrument sensitivity

## Slide 4: Automated Sample Prep
Advantages of Automated Sample Prep
  Rapid Turn Around Time: 30 to 45 Minutes for 6 Samples
  Cleaner Background Interferences: Closed Loop System
  Quality Results: Certified Pre-packaged Columns
  Green Technology: Lower solvent and power use
  QA/QC & Accreditation Requirements: Easier to Manage
  Computerized Method: Instrumentation based prep

## Slide 5: Sample Processing Workflow
Analysis of various matrices for PCDD/Fs and PCBs using extraction, cleanup and concentration.
Soxhlet extraction (typically up to 16-24 h).
Concentration step
Preparative multi-column chromatography involving various solvents and steps.
Can include acid-base-neutral silica, pure acidified silica, alumina, Florisil and carbon columns. Use of H2SO4 acid mixed with silica; NaOH mixed with silica Neutral Silica
Fractionation and Concentration

## Slide 6: Automation
Advantages of automated sample prep are:
Reduced time:
  Automated Pressurized Liquid Extraction (PLE) takes 60 min start-to-finish
  Manual Soxhlet up to 24 h.
Reduced cost: less labor involved, shorter turnover time per sample, less electricity use for PLE than Soxhlet.
Reduced volume: less solvent used.

## Slide 7: Automating the Workflow
Automated Pressurized Liquid Extraction (PLE) for sample extraction is fast (60 min), efficient (120 oC, 1500 psi), green (less power), reliable (long track record).
Users' choice of fully automated or semi automated clean up system: fast (30-45 min), low solvent usage

## Slide 8: Automated Cleanup
Fully automated sample load and elution.
Load Sample Extracts in hexane directly onto the system with no Manual Pretreatment
Easy to perform QC sample simultaneously with a Real sample.
  2 samples per module
Different column configuration: silica-carbon-alumina.
Uses no DCM, only Hexane and Toluene.
Total Clean Up time 30-45 min.
Low volumes

## Slide 9: Columns
Silica - PCB/PBDE-free Acid, Base and Neutral silica gel column (mini, classical, classical plus, high capacity, XL).
Alumina – PCB/PBDE-free basic alumina column.
Carbon – PCB/PBDE-free carbon/celite column.
Packed in disposable Teflon tubes; individually sealed in Mylar packaging; production in clean room environment.

## Slide 10
Classical Kit
Classical Plus Kit
Mini Kit
High-Capacity Kit
XL Capacity Kit
Column Kits

## Slide 11: SuperVap Concentrator

## Slide 12: SuperVap Evaporation
System pre-heated to 45-60 ºC.
Samples evaporated at stable T under 8 psi nitrogen.
1 mL extract vial transferred to GC vial (can have direct-to-vial feature).
Recovery standards added (nonane/dodecane).
Extract taken to 10 uL volume with a gentle stream of nitrogen at ambient temperature.

## Slide 13
GC vial
Glass Evaporation tube

## Slide 14: Using the GC-MS/MS
Agilent 7010

## Slide 15: GC-MS/MS Precision
Chart: 2-20 ppb cal std
| 2378-T4CDF | 2378-T4CDD | 12378-P5CDF | 23478-P5CDF | 12378-P5CDD | 123478-H6CDF | 123678-H6CDF | 234678-H6CDF | 123478-H6CDD | 123678-H6CDD | 123789-H6CDD | 123789-H6CDF | 1234678-H7CDF | 1234678-H7CDD | 1234789-H7CDF | OCDD | OCDF
QQQ-1117 | 1.8783 | 1.9711 | 9.7106 | 9.6539 | 9.8788 | 10.2861 | 10.6351 | 10.7549 | 11.055 | 8.002 | 8.8861 | 10.4964 | 9.4838 | 10.0437 | 9.7447 | 19.6181 | 20.1887
QQQ-1118 | 1.9043 | 2.0817 | 9.6814 | 10.4939 | 10.3053 | 10.7892 | 9.6284 | 10.1499 | 11.2233 | 8.3951 | 8.3814 | 10.341 | 9.7911 | 10.5242 | 9.7979 | 18.829 | 19.7513
QQQ-1119 | 1.7904 | 2.0922 | 9.7431 | 9.005 | 9.653 | 11.0093 | 9.571 | 10.19 | 11.3302 | 9.2575 | 8.9553 | 9.8488 | 10.0058 | 10.292 | 9.5036 | 18.6775 | 21.0205
QQQ-1120 | 1.8547 | 2.0311 | 9.6027 | 9.8353 | 9.2246 | 10.9459 | 10.8422 | 10.3957 | 10.1644 | 11.6917 | 9.8342 | 10.1282 | 9.5961 | 10.2579 | 9.5684 | 19.794 | 20.3478
QQQ-1121 | 1.8446 | 2.1562 | 10.1598 | 9.6978 | 9.8727 | 10.8081 | 9.9508 | 10.1725 | 9.3196 | 9.4802 | 7.7684 | 9.3288 | 10.2005 | 10.6166 | 9.8109 | 21.2838 | 20.2581
QQQ-1122 | 1.7863 | 1.8835 | 10.212 | 10.3798 | 9.663 | 11.3776 | 9.5735 | 8.8569 | 11.9018 | 9.4427 | 9.466 | 9.1969 | 9.3536 | 9.9591 | 9.3617 | 20.8142 | 20.4514

## Slide 16: GC-MS/MS Limit of Detection
Chart: Limit-of-detection
| 2378-T4CDF | 2378-T4CDD | 12378-P5CDF | 23478-P5CDF | 12378-P5CDD | 123478-H6CDF | 123678-H6CDF | 234678-H6CDF | 123478-H6CDD | 123678-H6CDD | 123789-H6CDD | 123789-H6CDF | 1234678-H7CDF | 1234678-H7CDD | 1234789-H7CDF | OCDD | OCDF
 | 0.035589 | 0.035041 | 0.14249 | 0.17494 | 0.10231 | 0.169 | 0.13324 | 0.15645 | 0.21027 | 0.17565 | 0.089233 | 0.20886 | 0.078718 | 0.15342 | 0.19045 | 0.091112 | 0.11684

## Slide 17: EP-110 13C Recoveries DD/F Matrices

## Slide 18: EP-110 13C PCBs Recoveries Matrices

## Slide 19: Semi-Automated Sample Cleanup
Simple to run, no computerized instrumentation
Fast: 30 - 45 min
Closed loop system to give a clean background, low level detection
Use certified pre-packaged columns
Green technology, only vacuum pump uses power
Economical column kits, choice of low fat and high fat column kits
No electronics or mechanical equipment to fail

## Slide 20: Characteristics of Semi-Automated System (EZPrep)
Small Solvent Reservoirs
Large Solvent Reservoirs
Turn Table
Stage 1 Manifold (waste)
3 Way Valve
Stage 2 Manifold (collection)
Large Solvent
Reservoir 100/180 mL
Small Solvent Reservoirs (40 mL)

## Slide 21: Stage 1: to waste
Vacuum pump
Waste
Rotatory workstation
Solvents
Column
assembly
Vacuum gauge

## Slide 22: Stage 2: collect
Manifold with collection tubes
Columns ready for elution

## Slide 23: Attributes
  Closed loop system:
    Eliminates background contaminants
    No washing needed.
    Capped solvent reservoirs
  Optimized for solvent reduction while obtaining highest possible recoveries
  Uses Hexane and Toluene, no Dichloromethane
  Easy sample loading on top of silica column via syringe vial
  Columns connect easy with SNAP connections

## Slide 24: Procedure
  Order of column assembly: Acid Silica-Carbon-Alumina
  Conditioning all columns with hexane (waste)
  Load sample
  Elute columns with hexane and transfer all target compounds to carbon and alumina (waste)
  Discard silica columns and remove carbon and alumina columns
  Elute carbon and alumina individually with toluene (reverse)
  Collect carbon fraction (PCDD/F and co-planary PCBs) and alumina fraction (mono- and di-ortho PCBs)

## Slide 25: 13C PCBs Recoveries EZPrep (%)

## Slide 26: 13C PCDD/F Recoveries EZPrep (%)

## Slide 27: Comparison of Native Data with Automated System

## Slide 28: 13C PCBs recoveries in soil
Chart: Recoveries (%)
| PCB_28 | PCB_52 | PCB_77 | PCB_81 | PCB_101 | PCB_105 | PCB_114 | PCB_118 | PCB_123 | PCB_126 | PCB_138 | PCB_153 | PCB_156 | PCB_157 | PCB_167 | PCB_169 | PCB_170 | PCB_180 | PCB_189 | PCB_209
 | 82.5 | 81.1667 | 82.1667 | 80.1667 | 79.3333 | 84 | 79.1667 | 82.3333 | 84 | 82.8333 | 86.8333 | 87.8333 | 86.1667 | 84 | 82.8333 | 78.8333 | 75.5 | 79.3333 | 83.5 | 86
10 g soil in toluene, n=6

## Slide 29: Conclusions (1)
Combination of PLE, SuperVap and EP-110/EZPrep give flexibility to (semi-) automate extraction, concentration and cleanup for a wide variety of compounds.
PLE (Pressurized Liquid Extraction)
  Reduced time:
    Automated Pressurized Liquid Extraction (PLE) takes 60 min start-to-finish
    Manual Soxhlet up to 24 h.
    Up to 8 samples in Parallel
  Reduced cost: less labor involved, shorter turnover time per sample, less electricity use for PLE than Soxhlet.
  Reduced volume: less solvent used.
EP110
  Uses No DCM
  Low solvent use 100-250 mL
  Total time from sample till data between 3-4.5 h
  Clean up step time between 30 and 45 min
  Modular 2 to 8 sample configurations in parallel
PLE and EP-110 can be configured in one system

## Slide 30: Conclusions (2)
Ezprep Semi-Automated Sample Cleanup
  Is a low solvent and fastest clean up system for Dioxins and PCBs (30 - 45 min)
High sample throughput → 18 samples/hour
  6 samples in parallel per station
  3 stations fit in one hood
System gives excellent recoveries for PCDD/F, PCB and PBDEs comparable to automated systems
Use of certified pre-packaged columns guarantees low native background

## Slide 31: Agilent 7010 MS/MS Dioxin Analysis
Provides a complete Dioxin/PCB workflow with FMS Sample Prep Systems
Easy to Use and Implement
Cost Effective Alternative to High-Res Mass Spec
  Lower instrumental purchase cost
  Lower cost of real estate
  Lower operational cost from lower training requirements and lower maintenance costs

## Slide 32: Questions?
r.addink@fms-inc.com